A Portable LAMP–CRISPR Lab-on-Chip Platform with Embedded Thermal Control and Optical Signal Processing: An End-to-End Theoretical Design

Om Krunal Awthankar

Abstract—Nucleic-acid tests can identify a predefined biological target with high molecular specificity, but conventional workflows commonly depend on centralized laboratories, trained personnel, thermal cyclers, benchtop optical instruments, and multiple manual liquid-handling steps. This paper presents an end-to-end theoretical design for a portable, programmable, and eventually sample-in–answer-out lab-on-chip platform that integrates loop-mediated isothermal amplification (LAMP), reverse-transcription LAMP (RT-LAMP), CRISPR-Cas12a sequence confirmation, embedded thermal regulation, microfluidic reaction handling, fluorescence detection, and edge-based result classification. The proposed device is not a generic “DNA-present” sensor. It is intended to answer whether a selected DNA or RNA sequence is present in a sample, with target programmability achieved by changing the LAMP primer set and CRISPR RNA.

The platform is divided into a reusable electronic reader and a sealed disposable cartridge. The cartridge concept contains zones for sample introduction, lysis and nucleic-acid preparation, isothermal amplification, controlled amplicon transfer, Cas12a reporter cleavage, optical interrogation, and waste containment. A staged engineering strategy is proposed: a tube-based proof of principle, a millifluidic sealed cartridge, and finally a microfluidic lab-on-chip with stored reagents and passive or mechanically actuated valves. The reader employs a microcontroller, closed-loop resistive heating, temperature sensing, light-emitting-diode excitation, spectral filtering, a photodiode transimpedance front end, and signal-processing firmware. Mathematical models for thermal behavior, proportional-integral-derivative control, photo-diode response, baseline compensation, and test classification are formulated. A validation framework covering analytical sensitivity, specificity, limit of detection, repeatability, contami-nation, cartridge integrity, power consumption, and usability is also defined. The paper further evaluates deployment use cases according to the operational friction removed by decentralized testing, identifying crop-pathogen screening, water surveillance, and food-processing quality control as particularly credible early applications. The outcome is a technically grounded project roadmap and architecture that can be implemented incrementally without presenting unverified clinical claims as experimental results.

Index Terms—CRISPR-Cas12a, DNA detection, embedded biosensor, fluorescence sensing, lab-on-chip, LAMP, microfluidics, molecular diagnostics, point-of-care testing, RT-LAMP.

M

Introduction

OLECULAR diagnostics identify biological targets by detecting characteristic nucleic-acid sequences. Unlike

a general fluorescence assay that merely indicates that DNA is present, a sequence-selective molecular test is designed to answer a much narrower question: whether a predefined genetic region associated with an organism, trait, mutation, or

biomarker is present in the tested sample. The identity of the target is encoded in molecular recognition elements, such as amplification primers and CRISPR RNA (crRNA), rather than in the electronic reader itself.

Polymerase chain reaction (PCR) is a foundational nucleic-acid amplification technology, but conventional PCR instru-ments require repeated and accurately controlled temperature cycling. Loop-mediated isothermal amplification (LAMP), introduced by Notomi et al., instead uses a strand-displacing polymerase and a set of target-specific primers to amplify DNA under approximately constant-temperature conditions [1]. Loop primers can further accelerate the reaction [2]. For RNA targets, reverse transcription can be coupled to LAMP, producing complementary DNA that is subsequently amplified. Isothermal operation reduces the thermal complexity of a portable reader, but amplification alone can still be affected by contamination, primer interactions, nonspecific products, and subjective endpoint interpretation.

CRISPR-associated nucleases have enabled an additional sequence-confirmation layer. Target binding by Cas12a acti-vates indiscriminate cleavage of nearby single-stranded DNA reporters, a property demonstrated by Chen et al. and used in the DNA Endonuclease-Targeted CRISPR Trans Reporter (DE-TECTR) concept [3]. A reporter containing a fluorophore and quencher is initially dim; Cas12a-mediated cleavage separates them and increases fluorescence. CRISPR diagnostics have subsequently been demonstrated with portable readouts and amplification-assisted workflows [4], [5]. The two recognition layers are complementary: LAMP enriches a selected sequence, while Cas12a and its guide provide a second molecular check. The long-term goal of this project is not only to build a heater and fluorescence reader, but to integrate the operations normally distributed across a laboratory into a compact cartridge-and-reader platform. A genuine lab-on-chip must address sample handling, lysis, nucleic-acid preparation, reagent storage, amplification, molecular confirmation, signal detection, controls, waste isolation, and user interaction. Micro-total-analysis systems and integrated nanoliter DNA-analysis devices established the broader concept of miniaturizing multi-step analytical workflows [6], [7]. Modern microfluidic architectures, including sliding-layer chips, demonstrate how fluid routing, partitioning, and multi-step reactions can be performed without

a network of external pumps [8], [9].

This paper therefore develops a complete theoretical ar-

chitecture and project plan for a portable LAMP–CRISPR lab-on-chip system. It makes no claim of clinical accuracy or experimental performance. Instead, it defines the intended function, system requirements, biological workflow, cartridge architecture, electronics, firmware, signal-processing methods, development phases, validation experiments, risks, limitations, and expected outcomes required to convert the concept into a defensible research prototype.

A. Principal Contributions

The main contributions of this work are:

- an end-to-end architecture separating a reusable embedded reader from a sealed disposable assay cartridge;

- a staged path from pre-extracted tube assays to a sample-in–answer-out microfluidic lab-on-chip;

- a two-stage LAMP/RT-LAMP and Cas12a workflow that avoids prematurely assuming that a robust one-pot chemistry is available;

- theoretical thermal, optical, electrical, and decision models for the reader;

- a firmware state machine with control-validity logic for positive, negative, invalid, and indeterminate results;

- a verification and validation framework that distinguishes engineering verification from analytical and clinical valida-tion;

- an explicit provenance and limitation analysis for the mathematical models and proposed decision metrics;

- an application analysis based on removal of sample-transport, laboratory-queue, and result-return friction; and

- a multidisciplinary work breakdown spanning biotechnology, electronics, computer science, mechanical design, and microfluidics.

- Project Definition, Aim, and Scope

  - Problem Statement

Existing nucleic-acid testing systems offer strong analytical capabilities but often depend on costly equipment, controlled laboratory environments, multiple transfer steps, and trained operators. Low-resource and field settings require a compact platform that can execute the assay with reduced manual han-dling while preserving target specificity, contamination control, traceability, and interpretable output. Integrating isothermal amplification and CRISPR detection into a portable device is challenging because biological reactions, thermal conditions, fluid routing, optical sensing, electronics, software, and quality controls must operate as one system.

  - Aim

The project aims to design and progressively prototype a portable molecular-testing platform that determines whether a predefined DNA or RNA target sequence is present in a sample. The intended final system consists of a reusable electronic reader and a target-specific, sealed, disposable lab-on-chip cartridge.

  - Research Question

The central research question is:

How can sample preparation, LAMP or RT-LAMP amplification, CRISPR-Cas12a sequence confirma-tion, microfluidic handling, embedded thermal regu-lation, fluorescence sensing, and digital classification be integrated into a portable and reprogrammable molecular-detection platform?

  - Specific Objectives

The project objectives are to:

    - define a safe, non-clinical model target for prototype development;

    - design a two-stage amplification-and-confirmation assay architecture;

    - develop a thermally controlled reaction platform suitable for LAMP and Cas12a incubation;

    - develop an optical subsystem for quantitative fluorescence acquisition;

    - design	a	sealed	cartridge	with	fluid-routing	and contamination-isolation features;

    - implement firmware for heating, timing, sensing, quality control, and result reporting;

    - implement signal processing for baseline compensation, filtering, thresholding, and test-validity checks;

    - characterize electronics, thermal stability, optical sensitivity, and power consumption;

    - define analytical validation experiments using synthetic or non-pathogenic targets; and

    - establish a route toward reagent storage, multiplexing, mobile connectivity, and manufacturing.

  - Non-Objectives and Boundaries

The initial project will not diagnose disease, replace lab-oratory PCR, process unapproved human samples, culture pathogenic organisms, or claim regulatory-grade performance. Clinical use requires ethics approval, biosafety governance, statistically powered testing, manufacturing controls, and regulatory authorization beyond the scope of an undergraduate prototype. The first implementation may use pre-extracted synthetic nucleic acid; this is a deliberate engineering stage rather than a claim of full sample-in–answer-out operation.

- Scientific and Engineering Background

- Target-Specific Detection Rather Than Generic DNA Detec-tion

Most environmental and biological samples contain nucleic acids. A device that only reports total DNA fluorescence would therefore have limited diagnostic value. Target-specific detection is achieved by selecting molecular recognition ele-ments that correspond to a unique or sufficiently discriminative region. In the proposed architecture, the LAMP primers and the Cas12a crRNA are changed when the target changes. The reader hardware remains substantially unchanged, although reaction temperature, timing, signal thresholds, and cartridge reagents may require assay-specific configuration.

- LAMP and RT-LAMP

LAMP uses inner and outer primers recognizing multiple tar-get regions and a strand-displacing DNA polymerase, enabling rapid amplification without repeated denaturation and annealing cycles [1]. Optional loop primers can increase reaction speed [2]. A simplified conceptual reaction is

(1)

time

T + P + E + N −−−−−−→ A,

isothermal

where T is the target template, P is the primer set, E represents enzymes and buffers, N denotes nucleotides, and A is the amplified product. Equation (1) is not a kinetic model; it represents the functional dependency that amplification requires compatible target, primers, chemistry, and temperature.

For RNA targets, reverse transcriptase first synthesizes complementary DNA. RT-LAMP can therefore be represented as

RNA −R−→T  cDNA −L−A−−M−P→ A.	(2)

The hardware benefit is that both processes can be performed under a nearly constant thermal regime when a compatible master mix is selected. The assay, however, remains sensitive to primer design, inhibitors, sample matrix, reaction composition, evaporation, and contamination.

- CRISPR-Cas12a Molecular Confirmation

Cas12a forms a ribonucleoprotein complex with crRNA. Recognition of a complementary DNA target adjacent to a suitable protospacer-adjacent motif activates target cleavage and collateral cleavage of single-stranded DNA reporters [3]. For a fluorescence-quenched reporter,

FQ reporter −a−c−ti−va−t−ed−−C−as−1−2→a  F + Q,	(3)

where separation of fluorophore F and quencher Q increases measured emission. Amplification-assisted Cas12 detection has been used in portable diagnostic research, including DETECTR-style workflows [4].

The proposed project uses Cas12a as a molecular confir-mation stage, not as a gene-editing tool. No modification of organisms or human genomes is involved. Cas12a is used in vitro as a programmable biochemical sensor.

- Why a Two-Stage Assay Is the Baseline

A one-pot assay is attractive because it reduces fluid transfers and contamination opportunities. However, ampli-fication enzymes and CRISPR nucleases may have different optimal temperatures, buffer requirements, reaction kinetics, and interference mechanisms. Published research has demon-strated one-pot approaches, but these systems require assay-specific optimization and cannot be assumed to be universally compatible [10], [11]. Therefore, the baseline design uses two physically separated chambers:

  - an amplification chamber optimized for LAMP or RT-LAMP; and

  - a CRISPR chamber optimized for Cas12a reporter cleavage. The lab-on-chip performs a sealed internal transfer after amplification. A one-pot variant is treated as an advanced research branch after the two-stage system is validated.

- Lab-on-Chip and Microfluidic Integration

A lab-on-chip miniaturizes and integrates laboratory op-erations within a network of channels, chambers, interfaces, and sensors. Early micro-total-analysis systems established the goal of integrated analytical handling [6]; later devices integrated DNA preparation, amplification, separation, and detection at nanoliter scales [7]. Sliding microfluidic platforms demonstrate that pumps and active valves are not always required; relative motion between patterned layers can establish or disconnect fluidic paths [8]. Digital LAMP–CRISPR has also been demonstrated in a multistep SlipChip architecture [9].

For this project, “lab-on-chip” is defined functionally: the cartridge must integrate several laboratory operations and constrain them in a reproducible sequence. Merely placing a tube into a heater is a useful prototype but is not considered the final lab-on-chip.

- Proposed End-to-End System Architecture

Fig. 1 shows the proposed system-level workflow. The user introduces the sample into a disposable cartridge, inserts the cartridge into the reusable reader, and initiates the test. The cartridge performs fluidic and biochemical operations while the reader supplies energy, timing, sensing, and computation.

- Reusable Reader

The reader contains:

- embedded processor and memory;

- heater drivers and temperature sensors;

- optical excitation and fluorescence-detection electronics;

- cartridge-presence and lid-interlock sensing;

- battery, charging, and power-management circuitry;

- display, indicators, buttons, and optional wireless interface; and

- firmware for assay execution, logging, controls, and classifi-cation.

- Disposable Cartridge

The cartridge contains target-specific and single-use compo-nents:

- sample inlet and leak-resistant closure;

- lysis/pre-treatment chamber or interface;

- nucleic-acid capture and wash zone, where required;

- LAMP/RT-LAMP reagents and reaction chamber;

- fluidic isolation and transfer mechanism;

- Cas12a, crRNA, and reporter chamber;

- optical window aligned with the reader; and

- terminal waste chamber that remains sealed after use.

- Functional Requirements

Table I lists the principal functional requirements. Numerical values are engineering targets and must be revised after assay selection and experimental characterization.

Sample introduction

Lysis and pre-treatment

Nucleic-acid capture / cleanup

LAMP or RT-LAMP

Sealed, controlled	Cas12a + crRNA + amplicon transfer	 fluorescent reporter

Optical interrogation

Positive / Negative / Invalid / Indeterminate

Disposable sealed cartridge

Reusable reader: thermal actuation, optics, electronics, firmware, and UI

Fig. 1. End-to-end architecture of the proposed portable LAMP–CRISPR lab-on-chip platform. Early prototypes may omit on-cartridge lysis and extraction, but these stages are required for the final sample-in–answer-out system.

TABLE I

Tentative Functional and Engineering Requirements

Requirement	Tentative design target	Rationale	Verification method

Target specificity	Programmable by primer set and crRNA

Amplification temperature	Assay-dependent; nominally near

60 °C–67 °C

Thermal stability	≤ ±0.5 ◦C at the sensor; chamber uniformity characterized separately

CRISPR incubation	Separate temperature/time profile supported

Optical acquisition		LED-on/LED-off differential reading with dark compensation

Result classes	Positive, negative, invalid, indeterminate

Containment	Cartridge remains sealed after amplification

Reader must support multiple assays without redesigning the entire instrument

Compatible with common LAMP/RT-LAMP chemistries Reduces reaction-to-reaction thermal variation

Cas12a chemistry may differ from LAMP chemistry

Suppresses ambient light and detector offset

Binary output alone can conceal control failure

Amplified products are major contamination sources

Sequence analysis followed by synthetic-target and non-target panels

Calibrated reference probe and thermal mapping

Time-series logging and multi-point validation

Two-zone heater test or controlled cooling transition

Optical blank, fluorescent standards, and stray-light tests

Firmware unit tests and blinded control runs

Dye leak test, pressure test, and environmental swabs

Portability	Rechargeable standalone reader	Enables field and classroom use	Mass, size, power, and battery

endurance measurement

Data integrity	Time-stamped run log and assay configuration ID

Safety	Lid interlock, over-temperature shutdown, current protection

Supports traceability and debugging	File checksums, event logs, and export

verification

Protects users and hardware	Fault injection and thermal/electrical

safety tests

- Lab-on-Chip Cartridge Design

  - Development Philosophy

The cartridge is the most multidisciplinary and risk-sensitive part of the project. It must be chemically compatible, optically clear where required, thermally conductive at reaction zones, mechanically sealable, inexpensive, manufacturable, and re-sistant to nonspecific adsorption and evaporation. Instead of attempting a fully microfabricated cartridge immediately, the project follows three generations.

    - Generation 1: Tube-Based Proof of Principle: The first system uses commercial reaction tubes and pre-extracted synthetic nucleic acid. Its purpose is to validate thermal control, optical acquisition, assay timing, and classification. Manual transfer may be used only inside controlled laboratory practice. This generation is not called sample-in–answer-out.

    - Generation 2: Sealed Millifluidic Cartridge: The second system uses channels and chambers with dimensions that can be fabricated by precision 3-D printing, laser cutting, adhesive lamination, or CNC machining. A sliding layer, blister, burstable seal, wax valve, or mechanically actuated membrane transfers amplified material without opening the cartridge. Millifluidics is chosen because it is easier to manufacture and inspect than true microfluidics while still testing fluidic sequencing and containment.

    - Generation 3: Microfluidic Lab-on-Chip: The final re-search stage miniaturizes channels and chambers, incorporates stored or dried reagents, reduces dead volume, and integrates sample preparation. Fabrication options include cyclic olefin copolymer, cyclic olefin polymer, polymethyl methacrylate, glass, or polydimethylsiloxane for research prototypes. Material choice must consider enzyme adsorption, gas permeability, autofluorescence, solvent compatibility, thermal expansion, bonding, and mass-manufacturing route.

  - Conceptual Cartridge Layout

Fig. 2 shows a conceptual two-layer cartridge. It is inten-tionally technology-neutral: the same functional layout may be implemented with a sliding plate, rotating disk, blister-driven laminate, or valve-controlled channel network.

  - Sample Introduction and Lysis

The sample inlet must accept a defined volume and close irreversibly or semi-permanently. The final sample-preparation method depends on the intended application and sample matrix. Saliva, swab eluate, plant tissue, food homogenate, and water concentrate require different approaches. The initial prototype should use a safe synthetic target in clean buffer, followed by a benign matrix-spike study. More complex lysis or purification is introduced only after the core detection chain is stable.

Sample inlet

Lysis / conditioning

Capture / cleanup

CRISPR

chamber

Isolation / transfer valve

Amplification chamber

Sealed waste reservoir

Optical window aligned above CRISPR chamber

Fig. 2. Conceptual cartridge layout. Geometry and actuation method remain tentative until fluidic simulations and benchtop tests are completed.

Possible future on-cartridge preparation strategies include thermal lysis, chemical lysis, membrane capture, or magnetic-bead capture. The specific chemistry must be selected and val-idated by the biotechnology team under institutional biosafety supervision. The theoretical architecture reserves space, heating access, magnets or actuators, and waste volume without prescribing a universal protocol.

  - Amplification Chamber

The amplification chamber must provide:

- controlled reaction volume;

- uniform thermal contact;

- low evaporation and bubble formation;

- chemical compatibility with amplification reagents;

- minimal dead volume; and

- isolation from the downstream detection chamber until amplification is complete.

A thin chamber improves thermal response but may increase surface-area-related adsorption. The design should include a vapor barrier or sealed headspace and should be evaluated for bubble nucleation during heating.

  - Amplicon Transfer and Contamination Control

Opening post-amplification tubes can release high-copy-number products into the workspace. The cartridge therefore

  - CRISPR Detection Chamber

The chamber stores or receives Cas12a, crRNA, reporter, and buffer. For early prototypes, reagents may be prepared immediately before use. A mature cartridge should use stabi-lized or lyophilized reagents with a defined shelf-life and rehydration sequence. The chamber includes an optically transparent window, a controlled optical path length, and geometry that avoids bubbles in the sensing region.

  - Controls

A credible molecular test requires controls. The proposed cartridge family should support:

- a no-template control during development to identify con-tamination;

- a positive assay control to verify reagent and reader function;

- an internal process or amplification control to identify inhibition or failed sample preparation; and

- optional reference fluorescence for optical self-test.

A single-channel prototype may run controls in separate cartridges. A later multi-chamber cartridge can execute target and control reactions in parallel.

- Embedded Electronics and Instrumentation

- Processing Platform

An ESP32-S3-class microcontroller is a suitable development platform because it combines digital interfaces, timers, analog acquisition support, wireless connectivity, and adequate process-ing for control and signal analysis. The final product may use a custom printed circuit board and a lower-power microcontroller. A separate single-board computer is not necessary for the first reader unless advanced imaging or a rich local interface is required.

- Thermal Subsystem

The baseline heater is a resistive polyimide or printed heater coupled to an aluminum or copper thermal spreader. Power is switched through a logic-level MOSFET. Temperature is measured using a calibrated negative-temperature-coefficient thermistor, platinum resistance temperature detector, or pre-cision digital sensor. A thermal fuse and independent over-temperature comparator provide protection beyond software control.

A lumped first-order thermal-resistance–thermal-capacitance model, obtained from conservation of energy with an equivalent heat-loss resistance, is [12]

C  dT (t) = P (t) − T (t) − Ta(t) ,	(4)

performs internal, one-directional transfer. Candidate actuation

th  dt	h

Rth

mechanisms are:

  - mechanical sliding of two patterned layers;

  - rupture of a preloaded blister or frangible membrane;

  - thermally actuated wax valve;

  - capillary burst valve; or

  - reader-driven miniature plunger or magnetic actuator.

The selected mechanism should minimize aerosol formation

where Cth is effective thermal capacitance, Rth is thermal resistance to ambient, Ta is ambient temperature, and Ph is applied heater power. The model is used for heater sizing and initial controller design; measured step-response data are later used for system identification.

∫

The closed-loop command may use the standard parallel-form proportional-integral-derivative controller [13],

and prevent backflow from the CRISPR chamber to the amplification chamber.

u(t) = Kpe(t) + Ki

t

e(τ )dτ + Kd

0

de(t)

,	(5)

dt

where e(t) = Tset − T (t). In practice, derivative action may be filtered or omitted if sensor noise is significant. Anti-windup and output limiting are required because the heater command is bounded.

If the amplification and CRISPR stages require different temperatures, the reader can use either one thermal zone with a controlled transition or two zones. A two-zone reader reduces waiting time but increases size, cost, calibration effort, and cartridge alignment requirements.

- Optical Excitation and Detection

The optical subsystem includes:

  - an excitation LED matched to the reporter fluorophore;

  - an excitation filter to narrow LED output;

  - a defined illumination geometry;

  - an emission filter that attenuates excitation leakage;

  - a silicon photodiode or spectral sensor;

  - a transimpedance amplifier; and

  - an analog-to-digital converter.

A right-angle geometry can reduce direct excitation leakage, while an epi-fluorescence geometry can simplify cartridge access but requires stronger filtering. The enclosure must be light-tight. Optical components must be selected after the reporter chemistry is fixed.

Photodiode current is approximated using the standard responsivity relation, with additive dark and background-current terms [14],

IPD(t) = RλPopt(t) + Idark + Iambient,	(6)

where Rλ is responsivity at the emission wavelength. For the ideal low-frequency circuit, the transimpedance amplifier produces [15]

VTIA(t) = Vref − IPD(t)Rf ,	(7)

with feedback resistance Rf . Feedback capacitance is selected to stabilize the amplifier and limit noise bandwidth.

The reader alternates LED-off and LED-on measurements.

A differential signal is

S(t) = Von(t) − Voff (t),	(8)

which suppresses ambient light and offset. A normalized fluorescence metric is

S(t) − Sbaseline

Idle

Cartridge check

Preheat

CRISPR

incubation

Transfer

Amplify

Optical read

Classify

Report / log

Safe fault

Any failure

Fig. 3. Proposed firmware state machine. Any safety, sensor, cartridge, or control failure transfers execution to a safe-fault state.

Battery capacity is then selected with derating for conversion loss, temperature, aging, and reserve margin.

E. Model Provenance, Assumptions, and Validation Status

The equations in this paper do not originate from a single LAMP–CRISPR publication, and they do not all have the same evidentiary status. Table II separates conceptual biochemical notation, established engineering models, and processing rules proposed specifically for this architecture. This distinction prevents an illustrative equation from being mistaken for a validated predictive model.

No numerical parameter in these equations is presented as experimentally identified. The lumped thermal model assumes that the heater, spreader, cartridge interface, and reaction volume can be represented by a small number of approximately uniform temperatures; spatial gradients and contact variation will create model error. The ideal photodiode/TIA relation omits detector shot noise, amplifier input-current and voltage noise, LED drift, optical-filter leakage, photobleaching, para-sitic capacitance, ADC quantization, and cartridge-to-cartridge variation. These effects must be added to a later uncertainty and noise budget. The normalized fluorescence and four-state

classifier are therefore research hypotheses rather than universal

FN (t) =

max(ϵ, S

reference

,	(9)

)

CRISPR diagnostic formulae.

where ϵ avoids division by zero and the reference may be an internal optical standard or baseline scale factor.

- Power Architecture

∫

The reader uses a rechargeable battery pack, protected charger, fuel gauge, and regulated rails. The heater is typically the dominant load. Analog optical circuitry should be isolated from heater-switching noise through layout, grounding, filtering, timing, and possibly a separate low-noise regulator. Energy required for a test is estimated by

ttest

Etest =	(Pheater + Pelectronics + Pdisplay) dt.	(10)

0

F. Tentative Reader Components

Table III presents a tentative, non-vendor-locked component architecture.

- Firmware, Software, and Data Processing

- Firmware State Machine

The reader operates through the state machine shown in Fig. 3. State transitions require explicit completion and safety conditions; a timer alone is insufficient.

Each test record includes assay identifier, firmware version, cartridge lot, timestamps, temperature trace, optical trace,

TABLE II

Provenance and Intended Use of the Mathematical Expressions

Expression	Status	Scientific or engineering basis		Permitted interpretation in this work

LAMP, RT-LAMP, and

reporter-cleavage arrows

Conceptual biochemical notation	Mechanisms described in the LAMP and

Cas12a literature [1], [3]

Shows required process order and dependencies; it is not a

mass-action or kinetic yield model

Thermal model, Eq. (4)	Established lumped engineering

approximation

First-law energy balance represented by equivalent thermal capacitance and resistance [12]

Supports preliminary heater sizing and controller design; parameters and spatial-gradient error must be identified experimentally

PID law, Eq. (5)	Established feedback-control structure

Standard parallel-form PID control [13]	Provides a candidate controller;

gains, filtering, anti-windup, and stability must be tuned on the actual reader

Photodiode and TIA relations, Eqs. (6)–(7)

Established first-order optoelectronic relations

Photodiode responsivity, dark current, and transimpedance conversion [14], [15]

Supports component selection and dynamic-range estimates; complete design must include shot noise,

op-amp noise, capacitance, filter leakage, and ADC error

LED-on/off subtraction, Eq. (8)

Normalized fluorescence and decision rule, Eqs. (9) and (11)

Proposed measurement method	Synchronous differential acquisition

commonly used to reject offsets and slowly varying background

Proposed in this work	Baseline correction, reference scaling, control-gated threshold classification

A design hypothesis for ambient rejection; performance must be quantified with representative cartridges and illumination conditions

Defines the intended software logic only; the reference term, thresholds, uncertainty band, and control criteria must be learned and locked from validation data

Energy integral	Standard physical definition	Energy is the time integral of total

electrical power

Provides battery-budget estimation after measured power traces and conversion losses are available

control status, final classification, and fault flags. Personally identifiable information is not required for engineering devel-opment and should not be stored by default.

- Signal Processing Pipeline

The signal-processing chain is:

  - detector stabilization and dark reading;

  - paired LED-off/LED-on sampling;

  - outlier rejection and median or moving-average filtering;

  - baseline estimation from early samples;

  - normalization against baseline or reference signal;

  - extraction of endpoint intensity, slope, and optional time-to-threshold;

  - control-channel validation; and

  - decision and confidence assignment.



A simple endpoint decision can be written as

positive,	FN (tf ) ≥ θp ∧ C = 1,

negative,	FN (tf ) < θn ∧ C = 1,

control timing. Machine learning is optional and should only be added if it improves performance on an independently held-out dataset. A transparent rule-based classifier is preferable during early development because it is easier to audit and validate.

- User Interface

The first reader should show:

- cartridge detected / not detected;

- heating, amplification, transfer, CRISPR, and reading progress;

- remaining or elapsed time;

- positive, negative, invalid, indeterminate, or device-fault result;

- battery and temperature warnings; and

- a run identifier for data retrieval.

Mobile and cloud connectivity are deferred until the local reader is reliable. Wireless features should not become dependencies for core assay execution.

D =



indeterminate, θn ≤ FN (tf ) < θp ∧ C = 1,

invalid,	C = 0,

(11)

- Tentative End-to-End Operating Sequence

The intended user and device sequence is as follows.

where C indicates successful control behavior. The thresholds θn and θp cannot be chosen from intuition; they must be estimated from validation data and locked for a specific assay version.

More advanced models may combine endpoint intensity, reaction slope, temperature deviations, optical self-test, and

The sequence is conceptual and must be converted into institution-approved standard operating procedures by trained biotechnology personnel.

- Assay selection: The reader loads the cartridge assay identifier and corresponding temperature, timing, optical, and classification parameters.

TABLE III

Tentative Reader Component Architecture

and stores an engineering log. The cartridge is removed and discarded according to approved laboratory waste

Subsystem	Candidate component class

Controller	ESP32-S3 development board; later custom MCU PCB

Heater	Polyimide resistive heater or PCB heater

Temperature sensing	Calibrated NTC

thermistor or Pt100/Pt1000 RTD

Heater driver	Logic-level

N-MOSFET with current sensing

Excitation	Narrow-band LED matched to reporter

Detector	Silicon photodiode or spectral sensor

Analog front end	Low-bias-current

op-amp TIA

ADC	MCU ADC for early tests; external precision ADC if required

User interface	OLED/LCD, status LED, buzzer, buttons

Power	Protected

Li-ion/Li-poly pack,

Selection criteria

ADC/interface capability, timers, wireless, low-power modes

Power density, uniformity, geometry, voltage

Accuracy, response time, calibration stability

Low loss, safe gate drive, fault monitoring

Wavelength, output stability, thermal drift Responsivity, dark current, active area Input bias, noise, rail range, stability Resolution, noise, sampling rate, reference stability

Readability, power, cleaning, simplicity Endurance, safety, availability,

procedures without being reopened.

- Project Development Methodology

The project is structured to reduce integration risk. Biological chemistry, reader hardware, and fluidics are first validated as modules before being combined.

  - Multidisciplinary Work Distribution

The electronics workstream owns thermal design, analog sensing, power, printed circuit boards, electromagnetic and ther-mal noise mitigation, and enclosure integration. The computer-science workstream owns firmware architecture, data logging, signal processing, test automation, configuration management, and optional interfaces. The biotechnology workstream owns target selection, primer and crRNA rationale, assay controls, reagent compatibility, sample matrix strategy, biosafety, con-tamination prevention, and analytical validation. Fluidic and mechanical tasks are shared or supported by faculty and collaborators with microfabrication experience.

Cross-training is essential. Electronics members must under-stand why reaction temperature, evaporation, and contamination matter; biotechnology members must understand sensor drift,

Enclosure

charger, regulators

Light-tight thermally insulated housing

serviceability

Optical isolation, thermal safety, cartridge alignment

calibration, and noise; computer-science members must under-stand that a classification model cannot compensate for failed controls or invalid chemistry.

- Sample introduction: A measured sample is added to the inlet and the inlet is sealed.

- Cartridge verification: The reader checks cartridge pres-ence, orientation, lot information, and lid closure.

- Sample preparation: The cartridge performs the applicable lysis, conditioning, and nucleic-acid capture/cleanup oper-ations. Generation-1 prototypes begin with pre-extracted synthetic material.

- Amplification: The reader heats the amplification chamber and records the temperature profile. The process is aborted if temperature leaves validated limits for an unacceptable duration.

- Internal transfer: The reader or cartridge actuator opens a one-way path and delivers a controlled fraction of amplicon into the CRISPR chamber without exposing it to the environment.

- CRISPR incubation: The chamber is held at the assay-specific temperature for Cas12a target recognition and reporter cleavage.

- Optical reading: The LED and photodiode acquire dark-compensated fluorescence measurements. Optical self-test and control reaction are evaluated.

- Classification: The firmware assigns positive, negative, indeterminate, or invalid status according to locked assay criteria.

- Reporting and containment: The reader displays the result

- Verification and Validation Framework

  - Verification Levels

Four levels are separated:

    - Engineering verification: Does the reader meet thermal, optical, power, mechanical, and software requirements?

    - Analytical validation: Does the assay detect its intended target and reject non-targets under defined laboratory conditions?

    - Matrix validation: Does sample preparation and detection work in the intended sample matrix?

    - Clinical or field validation: Does the system perform adequately on representative real-world specimens and users? This requires approvals and is outside the initial project.

  - Thermal Verification

Thermal testing includes:

- sensor calibration against a traceable or calibrated reference;

- warm-up time and overshoot;

- steady-state error and drift;

- spatial temperature uniformity across the chamber;

- response to ambient temperature and battery voltage;

- repeated runs and unit-to-unit variation; and

- open-sensor, short-sensor, stalled-heater, and runaway-heater fault tests.

TABLE IV

Proposed Phase-Wise Development Plan

Phase	Primary focus	Major activities	Exit criteria

- Governance and target definition

Select safe model target; identify laboratory supervisor; complete biosafety and contamination training; define requirements and records

Approved project scope, safe target, responsibilities, and risk register

- Reader thermal platform	Build heater, sensor, MOSFET driver, thermal

block, and control firmware; characterize warm-up, stability, overshoot, and spatial uniformity

- Optical platform	Build LED, filters, photodiode/TIA, ADC, and

light-tight enclosure; test fluorescent standards and optical blanks

- Tube-based molecular proof	Run approved LAMP/RT-LAMP and Cas12a

reactions using synthetic or non-pathogenic targets under supervision

- Integrated two-stage reader	Automate thermal sequence, transfer timing

proxy, optical readout, logging, and classification

- Sealed millifluidic cartridge	Fabricate chambers and transfer mechanism;

conduct water/dye tests, leak tests, dead-volume and bubble studies

Repeatable temperature profile under dummy thermal loads

Stable, monotonic signal with acceptable repeatability and dark compensation

Distinguishable target, non-target, and no-template control signals

End-to-end controlled run with predefined control logic

Repeatable sealed transfer and optical alignment without leakage

- Biological cartridge integration

- Sample preparation integration

Integrate reagents and assay into sealed cartridge; optimize chamber materials and transfer volume Add application-specific lysis and nucleic-acid cleanup; evaluate matrix inhibition

Reproducible target/control classification in sealed operation

Demonstrated processing of safe spiked matrix with internal control

- Microfluidic miniaturization	Reduce volume and footprint; study stored

reagents, bonding, manufacturability, and multiplexing

Prototype lab-on-chip meeting revised engineering requirements

- Validation and documentation

Execute analytical, electrical, thermal, usability, and robustness test plans; freeze design version

Complete dataset, limitations, design files, source code, and publication-ready report

  - Optical Verification

Optical tests include:

- dark offset and noise;

- ambient-light rejection;

- response linearity using fluorescent standards;

- LED intensity and temperature drift;

- optical alignment tolerance;

- cartridge-window variation;

- detector saturation and minimum measurable change; and

- cross-talk if multiple chambers are introduced.

  - Analytical Validation

The assay-development team should define, under supervi-sion:

- blank samples and no-template controls;

- positive synthetic target controls;

- non-target and near-neighbor sequence controls;

- dilution series across the expected detection range;

- replicate count and acceptance criteria;

- inclusivity across target variants where applicable;

- cross-reactivity panel;

- repeatability within one day and reproducibility across days, operators, reagent lots, and readers; and

- interference studies for the intended sample matrix.

The limit of blank (LoB), limit of detection (LoD), and, if quantitative claims are pursued, limit of quantification (LoQ) must be estimated using an appropriate statistical design. The platform should initially be framed as qualitative unless a validated relationship between target concentration and signal is demonstrated.

  - Software Verification

Firmware and software verification includes unit tests for temperature conversion, filtering, thresholds, timers, state transitions, fault handling, and data storage. Hardware-in-the-loop tests inject simulated temperature and optical signals. Version control, reproducible builds, configuration checksums, and immutable assay profiles reduce accidental changes during validation.

  - Blinding and Data Integrity

Whenever practical, the operator reading the device should not know which samples are positive or negative. Raw tempera-ture and fluorescence data must be preserved. Excluding failed runs after seeing the result is prohibited unless the exclusion criterion was defined before testing. Negative findings and invalid runs are part of the engineering evidence and should be reported.

- Risk Analysis and Mitigation

The principal technical, biological, and project risks are summarized below.

- Carryover contamination (critical): Amplified material can create false positives across later runs. The baseline mitigation is sealed post-amplification transfer, a one-way workflow, dedicated preparation tools, no-template controls, and closed waste containment.

- Poor primer or crRNA design (critical): Inadequate sequence design can cause weak sensitivity, nonspecific amplification, or no signal. Mitigation includes established

design criteria, in-silico specificity checks, synthetic controls, and iterative assay optimization.

  - Sample-matrix inhibition (high): A functioning reader may still produce false negatives if the sample suppresses extraction or amplification. An internal process control, cleanup or dilution strategy, and matrix-spike experiments are required.

  - Thermal gradients (high): Uneven temperature can pro-duce cartridge-to-cartridge variation. A thermal spreader, controlled mechanical contact, multi-point sensing during de-velopment, and finite-element or empirical thermal mapping are proposed.

  - Optical leakage and autofluorescence (high): Excitation leakage, cartridge fluorescence, and ambient light can desta-bilize thresholds. Spectral filtering, a light-tight enclosure, low-autofluorescence materials, reference measurements, and paired dark readings are required.

  - Bubble formation or evaporation (high): Gas bubbles and volume loss can interrupt transfer and distort optical readings. Chamber geometry, vapor barriers, controlled heating, bubble traps, and leak testing should be incorporated.

  - One-pot incompatibility (medium): Combining LAMP and Cas12a prematurely may suppress one or both reactions. The two-stage architecture remains the baseline until indepen-dently validated modules exist.

  - Reagent instability (high): Enzymes and reporters may have limited shelf life. Early work should use controlled storage; lyophilization and stabilizers are later research tasks with qualified collaborators.

  - Overfitted classification (medium): A model may appear accurate on the development dataset while failing elsewhere. Thresholds should be locked, raw data preserved, and perfor-mance evaluated on independent samples using interpretable logic wherever possible.

  - Scope expansion (high): Premature work on cloud services, AI, multiplexing, and extreme miniaturization can prevent completion of the core assay-reader chain. Stage gates are therefore mandatory.

  - Regulatory misrepresentation (critical): An unvalidated prototype could be mistaken for a clinical product. The system must remain research-use-only, avoid diagnostic claims, and require institutional approval for real specimens.

- High-Impact Deployment Use Cases and

Friction Reduction

The strategic value of the proposed platform is not simply that a molecular reaction is portable. Its value is the possibility of moving a predefined test closer to the location at which a time-sensitive decision must be made. The conventional pathway often includes sample packing, transport, accession at a centralized laboratory, queueing, manual processing, interpretation, and return of the result. A field-deployable cartridge-reader system can remove parts of this chain, while

positive, legally significant, indeterminate, or high-risk results can still be referred for confirmatory testing.

Crop-pathogen and nursery screening. Similar visible symp-toms may arise from infection, nutrient stress, water stress, heat, or chemical injury, while samples may travel to distant laboratories as disease spreads. A target-specific field cartridge could support earlier isolation of suspect plants, screening of nursery lots, and referral of only positive or uncertain samples. FAO estimates that plant pests and diseases cause losses of up to 40% of global crop production [16]. Every crop–pathogen pair would still require validated sampling, extraction, primers, crRNA, and interpretation.

Drinking-water and outbreak surveillance. Remote sources, schools, hostels, and emergency sites may be sampled infre-quently because collection, transport, and centralized testing are slow or costly. Local screening could identify target-specific contamination hotspots sooner and support source closure, boil-water advice, sanitation response, and confirmatory sampling. Microbiologically contaminated water can transmit major en-teric diseases [17]. Nucleic-acid detection does not necessarily demonstrate organism viability or infectivity, and dilute samples may require filtration or concentration.

Food-processing environmental screening. Product, raw material, or production zones may remain on hold while swabs, rinse water, and food samples undergo laboratory testing. Near-line screening of selected pathogen markers could accelerate cleaning verification, locate suspect processing zones, and reduce unnecessary quarantine of negative lots. Unsafe food produces substantial global health and socioeconomic burden [18]. Low organism concentration and complex matrices may require pre-enrichment, statistically valid sampling, and legally mandated confirmation.

Rural and primary-care infectious-disease testing. Sample referral can require transport, laboratory queues, a return patient visit, and empirical treatment before the cause is known. A narrowly defined and clinically validated cartridge could enable same-visit molecular information for referral, isolation, or pathogen-specific management. This is a high-impact but high-regulation pathway requiring ethics approval, clinical trials, controlled manufacturing, trained users, and authorization; the proposed student prototype is not a clinical diagnostic.

Antimicrobial-resistance marker screening. Culture and susceptibility testing can delay evidence-based therapy, while centralized genotypic tests may be unavailable at peripheral facilities. A multiplex cartridge could report a pathogen marker together with selected resistance genes, supporting earlier triage and decentralized surveillance. WHO has identified gaps in equitable access to diagnostics for bacterial, fungal, and resistant pathogens [19], [20]. Genotype does not perfectly predict phenotype, so phenotypic susceptibility testing remains important.

Vector, veterinary, and One-Health surveillance. Mosquito pools, ticks, livestock, wildlife, and environmental specimens are often transported to regional laboratories, delaying geo-graphic mapping and control measures. District teams could screen pooled vectors or approved animal and environmental

samples for selected targets, improving early warning and prioritizing specimens for central confirmation. One Health links human, animal, plant, and environmental detection and response [21]. Pooling, prevalence estimation, and field work still require rigorous sampling and biosafety protocols.

A. Recommended First Translation Path

The most credible first deployment pathway for the student team is a generic reader validated with synthetic nucleic acid, followed by a cartridge targeting a non-human agricultural or environmental application. Crop-pathogen or nursery screening is especially attractive because the value of earlier isolation is clear, sample access may be simpler, and the project can demonstrate true field portability without immediately making human clinical claims. Water and food-processing screening are strong subsequent directions but add matrix concentration, viability, enrichment, and regulatory-method-comparison challenges.

The principal friction reduction should therefore be measured at workflow level, not only as reaction time. Relevant metrics include time from collection to actionable screening result, number of manual transfers, percentage of samples requiring central referral, operator training time, contamination rate, invalid-test rate, cold-chain requirement, and cost per decision. A device that shortens the biochemical reaction but retains complex off-chip extraction, transport, or interpretation has not yet achieved the intended operational impact.

- Expected Outcomes

The project is expected to produce the following outputs.

- Technical Outcomes

  - A reusable embedded reader capable of executing pro-grammable thermal profiles and fluorescence measurements.

  - A validated thermal-control module with logged temperature trace and fault protection.

  - A characterized optical detector capable of distinguishing reporter fluorescence from blank and background signals.

  - A sealed two-stage cartridge prototype that isolates amplifi-cation from Cas12a detection and waste.

  - A firmware platform implementing assay sequencing, con-trol checks, data logging, and four-state result reporting.

  - A documented route from pre-extracted samples to inte-grated lysis and nucleic-acid cleanup.

  - A target-reprogrammable architecture in which assay iden-tity is changed primarily through cartridge chemistry and digital assay configuration.

- Scientific Outcomes

The project should generate data on thermal behavior, fluores-cence detection limits, cartridge transfer repeatability, material compatibility, contamination control, and assay performance. These data can support a later experimental paper. The present theoretical paper provides the system hypothesis and validation framework; a subsequent paper must report measured results, uncertainty, failures, and limitations.

- Educational Outcomes

The project provides a shared platform for students to learn embedded control, analog instrumentation, microfluidics, molecular recognition, assay validation, data science, systems engineering, and research ethics. Its value lies not merely in assembling modules but in understanding how failure in one domain propagates through the complete system.

- Translation and Impact Outcomes

In addition to the reader, cartridge, firmware, and validation dataset, the project should produce an application-selection framework that compares candidate targets by unmet work-flow need, sample accessibility, biosafety, matrix complexity, confirmatory-testing pathway, regulatory burden, and measur-able friction reduction. Reprogrammability does not imply that a new assay becomes valid merely by changing primers and crRNA; every target requires its own chemistry, controls, thresholds, sample-preparation method, validation dataset, and risk assessment.

- Limitations and Open Research Questions

The design has several unresolved limitations. First, sample preparation may dominate complexity, especially for viscous, inhibitor-rich, or low-concentration samples. Second, a two-stage assay requires reliable transfer and may increase total test time. Third, microfluidic surfaces can adsorb proteins or nucleic acids and can introduce autofluorescence. Fourth, fluorescence thresholds depend on assay chemistry, optical geometry, cartridge variation, temperature, and reagent lot. Fifth, reagent stabilization and room-temperature storage are separate research challenges. Sixth, multiplexing creates primer interactions, optical cross-talk, and more complex controls. Seventh, a laboratory prototype does not establish clinical utility.

Important open questions include:

- Which cartridge material and bonding process provide the best balance of optical quality, chemical compatibility, and low-cost fabrication?

- Can a passive transfer mechanism deliver sufficiently consistent volume without pumps?

- Should the reader use one thermal zone, two thermal zones, or a thermally compatible Cas12 enzyme for future one-pot operation?

- What internal control best distinguishes failed extraction, amplification inhibition, optical failure, and reagent degra-dation?

- Can reagent stabilization achieve an acceptable shelf life under the intended storage conditions?

- What signal features provide robust qualitative classification across readers, cartridges, and ambient conditions?

- Ethics, Biosafety, and Responsible

Development

Early development should use synthetic nucleic-acid frag-ments, non-pathogenic organisms, or other institutionally approved materials. Work with human-derived samples requires

ethics approval and informed-consent procedures where appli-cable. Work with pathogens or cultured organisms requires the appropriate biosafety level, trained personnel, facilities, and waste handling. The device must be labeled for research use only until validated and authorized for a defined use.

The project should minimize collection of personal data. Genetic data can be sensitive, and future mobile or cloud features require data minimization, encryption, access control, retention policies, and applicable legal review. The device should not infer identity, ancestry, health status, or treatment decisions outside an approved and validated scope.

- Conclusion

This paper has presented an end-to-end theoretical design for a portable LAMP–CRISPR molecular-detection platform and its progression toward a lab-on-chip. The proposed system detects a predefined DNA or RNA sequence rather than generic nucleic-acid presence. LAMP or RT-LAMP supplies isothermal target amplification, while Cas12a provides a second programmable recognition layer and fluorescence-generating collateral reporter cleavage. The reusable reader integrates thermal control, optical sensing, power management, firmware, and local result inter-pretation. The disposable cartridge integrates sample handling, amplification, sealed transfer, CRISPR detection, controls, and waste containment.

The project is deliberately staged. A tube-based reader first proves thermal, optical, and assay principles; a sealed millifluidic cartridge then proves fluid handling and contami-nation control; and a final microfluidic cartridge integrates sample preparation and stored reagents. This progression converts an ambitious interdisciplinary concept into verifiable engineering increments. The immediate outcome is a research-ready architecture, requirements set, mathematical model, and validation plan. The long-term outcome is a modular platform that can be adapted to carefully validated applications while maintaining scientific, biosafety, ethical, and regulatory discipline.

Acknowledgment

The authors may acknowledge faculty mentors, laboratory staff, institutional facilities, and funding agencies here after receiving permission. This manuscript currently describes a theoretical design and does not claim experimental or clinical validation.

- J. S. Gootenberg et al., “Multiplexed and portable nucleic acid detection platform with Cas13, Cas12a, and Csm6,” Science, vol. 360, no. 6387,

pp. 439–444, 2018, doi: 10.1126/science.aaq0179.

- A. Manz, N. Graber, and H. M. Widmer, “Miniaturized total chemical analysis systems: A novel concept for chemical sensing,” Sensors and Actuators B: Chemical, vol. 1, nos. 1–6, pp. 244–248, 1990, doi: 10.1016/0925-4005(90)80209-I.

- M. A. Burns et al., “An integrated nanoliter DNA analysis device,” Science, vol. 282, no. 5388, pp. 484–487, 1998, doi: 10.1126/sci-ence.282.5388.484.

- W. Du, L. Li, K. P. Nichols, and R. F. Ismagilov, “SlipChip,” Lab on a Chip, vol. 9, no. 16, pp. 2286–2292, 2009, doi: 10.1039/B906537A.

- Z. Yu et al., “Parallel multistep digital analysis SlipChip demonstrated with the quantification of nucleic acid by digital LAMP-CRISPR,” Lab on a Chip, vol. 22, pp. 2954–2964, 2022, doi: 10.1039/D2LC00266C.

- Z. H. Cheng et al., “Tunable control of Cas12 activity promotes universal and rapid one-pot CRISPR diagnostics,” Nature Communications, vol. 16, 2025, doi: 10.1038/s41467-025-56516-3.

- C. R. H. Lynch et al., “Optimization and benchmarking of RT-LAMP-CRISPR-Cas12a assays for nucleic-acid detection,” 2025. [Online]. Avail-able: https://pmc.ncbi.nlm.nih.gov/articles/PMC11899638/. Accessed: Jul. 17, 2026.

- T. L. Bergman, A. S. Lavine, F. P. Incropera, and D. P. DeWitt, Fundamentals of Heat and Mass Transfer, 8th ed. Hoboken, NJ, USA: Wiley, 2017.

- K. J. A˚ stro¨m and T. Ha¨gglund, Advanced PID Control. Research Triangle Park, NC, USA: ISA, 2006.

- Hamamatsu Photonics K.K., “Technical note: Si photodiodes,” 2025. [Online]. Available: https://www.hamamatsu.com/content/dam/ hamamatsu-photonics/sites/documents/99SALESLIBRARY/ssd/si  pdkspd9001e.pdf. Accessed: Jul. 17, 2026.

- Texas Instruments, “1 MHz, single-supply, photodiode amplifier reference design,” TIDU535, Nov. 2014. [Online]. Available: https://www.ti.com/ lit/pdf/tidu535. Accessed: Jul. 17, 2026.

- Food and Agriculture Organization of the United Nations, “About FAO’s work on plant production and protection,” 2026. [Online]. Available: https://www.fao.org/plant-production-protection/about/en. Accessed: Jul. 17, 2026.

- World Health Organization, “Drinking-water,” Sep. 13, 2023. [Online]. Available: https://www.who.int/news-room/fact-sheets/detail/ drinking-water. Accessed: Jul. 17, 2026.

- World Health Organization, “Estimating the burden of foodborne diseases,” 2026. [Online]. Available: https://www.who.int/activities/ estimating-the-burden-of-foodborne-diseases. Accessed: Jul. 17, 2026.

- World Health Organization, “Antimicrobial Resistance Diagnostic Initiative,” Jun. 27, 2024. [Online]. Available: https://www.who.int/ publications/i/item/9789240072015. Accessed: Jul. 17, 2026.

- World Health Organization, “Landscape analysis of commercially available and pipeline diagnostics for bacterial priority pathogens and antimicrobial resistance,” 2025. [Online]. Available: https://www.who. int/publications/b/77189. Accessed: Jul. 17, 2026.

- World Health Organization, “One Health,” May 8, 2026. [Online]. Available: https://www.who.int/news-room/fact-sheets/detail/one-health. Accessed: Jul. 17, 2026.

References

  - T. Notomi, H. Okayama, H. Masubuchi, T. Yonekawa, K. Watanabe,

N. Amino, and T. Hase, “Loop-mediated isothermal amplification of DNA,” Nucleic Acids Research, vol. 28, no. 12, p. e63, 2000, doi: 10.1093/nar/28.12.e63.

  - K. Nagamine, T. Hase, and T. Notomi, “Accelerated reaction by loop-mediated isothermal amplification using loop primers,” Molec-ular and Cellular Probes, vol. 16, no. 3, pp. 223–229, 2002, doi: 10.1006/mcpr.2002.0415.

  - J. S. Chen et al., “CRISPR-Cas12a target binding unleashes indiscriminate single-stranded DNase activity,” Science, vol. 360, no. 6387, pp. 436–439, 2018, doi: 10.1126/science.aar6245.

  - J. P. Broughton et al., “CRISPR-Cas12-based detection of SARS-CoV-2,” Nature Biotechnology, vol. 38, pp. 870–874, 2020, doi: 10.1038/s41587-

020-0513-4.